## Slide 1: Eagle 2
www.rkiinstruments.com
Service Training Module
Notes: <number>

## Slide 2: Scope
MODES – normal mode, display mode, leak check, bar hole, inert mode, methane elimination (modes in blue are factory set to off unless requested upon order)
MENUS – Set up menu, Factory menu
SERVICE MAINTENANCE– Common issues, calibration (auto cal vs single cal), filter replacements, sensor replacement, pump replacement/rebuild, flow adjustment, Main board replacement, Display PCB replacement, PID cleaning.
www.rkiinstruments.com

## Slide 3
www.rkiinstruments.com
EAGLE 2 MODES
Normal Mode
Display Mode
Leak Check/Bar Hole
Notes: <number>

## Slide 4: Normal Mode
Press and Hold POWER ENTER until unit turns on. The unit will go into warm up and show some instrument parameters.
Perform fresh air calibration
Perform a flow integrity test (make sure unit goes into a “fail low flow” alarm) to make sure there is not leaks in the system.
www.rkiinstruments.com

## Slide 5: Display Mode
Peak reading (min and max, values will reset when instrument is turned off).
Battery info
Gases display – display all, scroll auto, scroll manual
Methane Elimination (can be off)
Relative response (can be off in set up mode)
STEL
TWA
View Alarm Settings
Select User ID
Select Station ID
Operation time
Date and Time
Data log remain
www.rkiinstruments.com

## Slide 6: Leak Check & Bar Hole Mode
If the leak check or bar hole mode is on in SET UP MODE MENU, you should see this mode select screen at start up:
NOTE: Bar hole mode will only show up in the mode select screen if you have a % vol TC methane sensor and LEL sensor installed, both configured to methane.
www.rkiinstruments.com

## Slide 7
www.rkiinstruments.com
EAGLE 2 MENUS
Set up Mode Menu
Factory Mode Menu
Notes: <number>

## Slide 8: Setup Mode Menu (UP+DOWN+POWER)
NOTE: The Eagle 2 is not a gas monitoring device when it is in Setup Mode.
To access Setup Mode, follow the steps below:
Start with the Eagle 2 OFF.
Press and hold the UP (AIR/YES) + DOWN (RANGE/SHIFT) buttons and then hold POWER ENTER. Release when you hear a beep.
An “S” on the lower right hand side will show up on the screen. The following list will be displayed six at a time:

## Slide 9
  Set Date & Time
  Set Date Format
  Set Battery Type – Alkaline or Ni-MH
  Configure Channels – Turn channels on or off/switch channel sequence
  Configure Gases – Usually used to change LEL gas name to another combustible gas
  Catalytic Units – Switch from ppm, LEL, % vol if catalytic sensor is installed.
  Relative Response – catalytic sensor response chart can be turned on to show up in unit display mode.
  Alarm Points – change and view alarm points for each gas
  Alarm Latching – self reset or latching option (factory set: latching)
  Alarm Silence – on/off. Allows user to silence audible alarm even if alarm is not completely cleared.
Setup Mode Menu (UP+DOWN+POWER)

## Slide 10: Setup Mode Menu (UP+DOWN+POWER)
  User/Station ID – On/Off in the unit display mode.
  Adj Autocal Values - Change calibration values to match cylinder.
  Backlight Delay
  Auto Fresh Air Adj – fresh air at unit turn on (factory set to off)
  Data Log Interval – 5 sec (10 days) - 10 min (2000 days)
  Data Log Overwrite – On/Off. Overwrites oldest data if on.
  Data Log Memory – On/Off for prompt to clear data memory.
  Adjust Contrast – Change display contrast
  Cal Reminder – On/Off (cal c-limit)
  Cal Past Due Act – Confirm to use, must cal, notification only
  Cal Interval – Amount of time between calibration reminder. Factory setting is 90 days.

## Slide 11: Setup Mode Menu (UP+DOWN+POWER)
  LC/BH Mode Select – Turns on/off leak check and bar hole modes.
  BH Measuring Time – Amount of time for each barhole measurement in sec. Factory setting is 30 secs.
  Zero Follower
  Zero Suppression
  Confirmation Alert – Beeps/lights up every 15 min to confirm unit is still functioning.
  Change Password – On/Off
  Default Settings – Restores to RKI standard configurations
  Lunch Break – On/off
  Span Factor – On/off. The range displayed after performing single calibration.
  Select Language
  Normal Operation

## Slide 12: Factory Mode Menu (UP+DOWN+DISPLAY+POWER)
To access Factory Mode, follow the steps below:
Start with the Eagle 2 OFF.
Press and hold the UP (AIR/YES) + DOWN (RANGE/SHIFT) + DISPLAY buttons and then hold POWER ENTER. Release when you hear a beep.
An “F” on the lower right hand side will show up on the screen.
ROM # and CHECKSUM will appear. Press DISPLAY/NO
mV screen will show up. Press POWER ENTER
Password: 1994, press POWER to continue
www.rkiinstruments.com

## Slide 13: Factory Mode Menu (UP+DOWN+DISPLAY+POWER)
In Factory mode you can press “AIR/YES” or “DISPLAY/NO” to answer the menu questions:
Enter or Change serial number?
Change Methane Elimination Voltage?
Enter user defined relative responses?
Set alternate range for standard sensors?
IR auto range change?
Eagle 2 logo display?
www.rkiinstruments.com

## Slide 14: Factory Mode Menu (UP+DOWN+DISPLAY+POWER)
Demand zero set?
Span factor display in single cal?
Methane elimination mark display in normal mode?
Symbol mark display in normal mode?
Data logging remain display in display mode?
Data logging remaining time
Active gases display in warm-up mode?
Alarm display in warm-up mode?
Save default instrument settings?
www.rkiinstruments.com

## Slide 15
www.rkiinstruments.com
EAGLE 2 SERVICE MAINTENANCE
Auto Calibration, Single Calibration, Common Issues, Filter replacement, sensor replacement, pump replacement/rebuild, Main board replacement, Display PCB replacement, PID cleaning
Notes: <number>

## Slide 16: Service Maintenance
Some maintenance procedures may be performed by the user. Most of these will be covered in this class.
More complicated procedures must be performed by RKI Instruments, Inc. (i.e. adding new sensors to the unit other than standard 4 gas)
For all maintenance procedures, be sure that the Eagle 2 is turned off and batteries are removed.

## Slide 17: Auto Calibration
NOTE: This describes a standard 4 gas unit LEL/O2/H2S/CO

## Slide 18: Auto Calibration
Required Equipment:
  Calibration cylinder – typically use 50% LEL methane, 12% O2, 25 ppm H2S, 50 ppm CO balance N2. O2 cal gas needs to be 15%vol and below.
  Demand flow regulator
  Non-absorbent tubing - P/N 06-1248RK-03
  NOTE: Before performing a span calibration with the cylinder bottle, perform a flow test and a fresh air adjustment first.

## Slide 19: Auto Calibration
Turn on the Eagle 2 unit by pressing the POWER ENTER button.
Once warm up is complete and the unit is in Normal Measuring mode, perform fresh air adjustment by pressing and holding the AIR button. The unit will tell you when to release the AIR button.
While in Normal Measuring Mode, Go into the calibration mode by pressing and holding the RANGE/SHIFT button, then press the DISPLAY/ADJUST button and release.
  Note: Backdoor password: 1196
The Calibration Mode Screen will be displayed with the cursor next to AUTO CALIBRATION.

## Slide 20: Auto Calibration
Press the POWER ENTER button to enter Auto Cal. This will allow you to calibrate all standard 4 sensors simultaneously.
The auto calibration values will be displayed.
  CAL GAS VALUES
  CH4 50% LEL
  OXY 12% vol.
  H2S 25.0 ppm
  CO 50 ppm
  ENTER TO BEGIN CAL
  NOTE: Make sure that these values match your cylinder, as these are the values that the Eagle 2 will adjust the calibration to.

## Slide 21: Auto Calibration
If the gas concentrations to be used are different than what is displayed, press and hold the RANGE/SHIFT button then press the DISPLAY button. The LCD will indicate the following:

## Slide 22: Auto Calibration
Use the AIR YES or RANGE SHIFT buttons to move the cursor to the channel you wish to change and press ENTER.
The current value will begin to blink
Use the AIR YES or RANGE SHIFT buttons to increase or decrease that value, respectively.
To confirm the new calibration value, press ENTER.
Follow the same procedure for changing other values.
Make sure that the calibration values on the screen match the calibration values on the calibration cylinder.

## Slide 23: Auto Calibration
To return to the auto calibration screen, scroll down to END and press ENTER.
The screen will then ask if you would like to save the changes. Press YES to save them.
You will be returned to the auto calibration screen and the new calibration values will be displayed.
Press the ENTER button to start the calibration.
Attach the demand flow regulator onto the calibration cylinder.
Connect the tubing from the demand flow regulator to the rigid tube on the probe.

## Slide 24: Auto Calibration
Allow readings to stabilize or a maximum of 2 minutes
Press the POWER ENTER RESET button to complete calibration
Remove calibration gas from probe
If all channels passed calibration, the display will prompt you to remove the gas and then it will return to the Calibration Mode Screen.
If any channels failed to calibrate, the display will show which sensors failed and the buzzer and alarm LED’s activate. Press ENTER to return to the Calibration Mode Screen.

## Slide 25: Auto Calibration
To return to Normal Operation, use RANGE SHIFT to place the cursor next to NORMAL OPERATION and press ENTER.
Remove the regulator from the calibration cylinder.
Store the calibration materials in a safe place.

## Slide 26: Single Calibration
When do we use Single Calibration?
Single calibration is used to calibrate each sensor individually.
To check the sensor life using the span value shown after single calibration. (NOTE: the wider the range of the span, the better the sensor).
When a sensor fails auto calibration.

## Slide 27: Single Calibration
Enter Calibration Mode by pressing and holding RANGE SHIFT, then pressing DISPLAY and releasing both.
Use the RANGE SHIFT button to move the cursor next to SINGLE CALIBRATION and press ENTER.
Attach the demand flow regulator to the calibration cylinder and connect tubing.
Insert the probe into the Eagle 2.

## Slide 28: Single Calibration
Select sensor to calibrate:
>ESCAPE
CH4 OXY
H2S CO
Use the RANGE SHIFT and AIR YES buttons to move the cursor next to the channel you wish to calibrate and press ENTER.
Use the RANGE SHIFT and AIR YES buttons to adjust the calibration value, if necessary.

## Slide 29: Single Calibration
Use the RANGE SHIFT and AIR YES buttons to adjust the calibration value to match your cylinder, if necessary.
Press ENTER to begin calibration; the screen will flash CAL IN PROGRESS.
Connect the sample tubing to the probe and allow reading to stabilize.
Press ENTER to finish the calibration.
If the calibration was successful, the Eagle 2 will display the span value and then return to the sensor selection screen.

## Slide 30: Single Calibration
If the calibration fails, the display will show the sensor failure and the span. The buzzer and LEDs will pulse.
Press ENTER to silence the alarm and return to the sensor selection screen.
Repeat the process for any other sensors.
When finished, move the cursor to ESCAPE and press ENTER to return to the calibration screen.
Move the cursor to NORMAL OPERATION and press ENTER.

## Slide 31
Take the Eagle 2 to a fresh-air environment.
Attach probe to the inlet at the front of the instrument. Make sure the filters are all clean on the probe and on the inside of the unit before performing this test.
Turn the Eagle 2 on and allow the instrument to warm up.
Place your finger over the end of the probe and verify that the Eagle 2 indicates a ”FAIL LOW FLOW” alarm. This verifies that the probe and inlet are leak tight.
Press the RESET switch to restart pump.
NOTE: If the Eagle 2 does not go into low flow alarm, check for leaks in the probe, internal tubings, and filters. If these look good, a flow adjustment procedure to reset the differential pressure (DP) switch needs to be performed. See next slide on how to set the DP switch.
Flow Integrity Test

## Slide 32: Setting the DP Switch
Required equipment:
  Small flat head screw driver,
  0-2 SCFH flow meter (RKI P/N 31-0023RK)
  Male quick connect fitting (RKI P/N 17-0505RK)

## Slide 33: Setting the DP Switch
The DP (differential pressure) switch monitors the flow pressure in the flow system.
Normal flow for the Eagle 2 is between 1.6 and 2 SCFH, depending on the number of sensors.
Trip point for the DP switch is .6 SCFH +/- .1.
DP switch adjustment screw

## Slide 34: Setting the DP Switch
Before setting the DP switch, check the following first:
Ensure that the filters are clean as plugged filters can cause the pump flow to decrease and activate the ”fail low flow” alarm.
Make sure the tubings are clean and that there is no leak on the inside of the system.
NOTE: The goal is to make sure the unit will activate the flow fault alarm and shut down the pump when the flow rate is reduced to 0.6 SCFH +/- 0.1 SCFH.

## Slide 35: Setting the DP Switch
Turn on the Eagle 2 and allow unit to warm up.
Perform a fresh air adjust, if needed.
Attach the flow meter to the inlet fitting of the Eagle 2 making sure that the flow meter valve is fully open.
Verify that the flow rate is between 1.6 and 2 SCFH.

## Slide 36: Setting the DP Switch
Adjust the flow meter valve to 0.6 SCFH+/-0.1.
Slowly adjust the DP switch pot, as needed, to activate flow fault alarm at 0.6 SCFH +/- 0.1. (Turning the adjustment ccw raises trip point, cw lowers trip point)

## Slide 37: Replacing Batteries
To access the battery compartment, unscrew the thumbscrew located at the back of the unit and pull the battery case away from the bottom case.
Remove the old batteries and make sure the battery compartment is clean.
Carefully put the new batteries into the battery compartment being sure to follow the diagram and push the batteries in all the way.
Reinstall the battery case and tighten thumbscrew.
After replacing Ni-MH batteries, charge batteries for 4 hours for a complete charge.
NOTE: Fresh alkaline batteries will operate the Eagle 2 for 18 hours. Charged Ni-MH batteries will operate the Eagle 2 for 20 hours at 25C (77F).

## Slide 38: Replacing Hydrophobic Filter
To access the filter, turn the Eagle 2 upside down and unscrew the 3 screws holding the top case to the bottom case.
Carefully lift the top case off and place next to the bottom case.
Pull the gray connecting tubes off each end of the filter.
Install the new filter with the red RKI logo facing the front of the Eagle 2.
Hydrophobic Filter
NOTE: Replace the internal hydrophobic filter when it becomes dirty or clogged. Internal filter P/N 33-0173RK.

## Slide 39: Replacing Charcoal Filter
Grasp the charcoal filter end cap and pull it off the charcoal filter.
Grasp the top of the charcoal filter firmly and pull it out of the flow block. (be careful as contents might spill)
Insert the new charcoal filter into its position in the flow block and push it in until it bottoms out.
Insert the charcoal filter end cap into the end of the new charcoal filter. (Charcoal filter P/N 33-6090RK)
NOTE: Replace filter if exposed to high levels of HC’s, when the CO sensor is replaced, or if the CO sensor responds to H2S.
Charcoal Filter

## Slide 40: Replacing a Sensor
CO
Sensor
red and black wire
Oxygen Sensor
LEL Sensor red, white, green, black wire
H2S
Sensor
orange and black wire
NOTE : Always recalibrate your instrument after replacing sensors. If you are getting unstable readings, let the Eagle 2 warm up for about 10-15 minutes before recalibrating.

## Slide 41: Replacing LEL Sensor
Unscrew and remove the 2 screws holding the LEL sensor retainer.
Grasp the LEL sensor connector and gently pull it up until it either disengages from the LEL sensor or the LEL sensor comes out of the flow block.
If the sensor came out, remove the sensor from the connector. If the sensor stayed in the flow block, pull it up and remove it.
Connect the new LEL sensor and the sensor connector.
Insert the LEL sensor into the sensor chamber.
Line up the LEL sensor retaining bracket and screw it back into place.
Calibrate the Eagle 2 after replacing the LEL sensor.

## Slide 42: Replacing Oxygen Sensor
Unscrew the 2 screws that hold the oxygen sensor bracket and remove the bracket. Note the routing of the oxygen sensor cable to the main PCB.
Move the hydrophobic filter towards the side wall and pull the oxygen sensor out of the flow block.
Grasp the connector on the end of the sensor cable and pull it away from the main PCB.
Insert the new oxygen sensor face down into the flow block.
Route the sensor cable and connect it to the main PCB.
Reinstall the oxygen sensor bracket.
Calibrate the Eagle 2 after replacing the O2 sensor.

## Slide 43: Replacing CO or H2S Sensor
Grasp the sensor firmly and rock it back and forth while pulling straight up.
Grasp the connector on the end of the sensor cable and pull it away from the main PCB.
If sensor becomes difficult to remove from the chamber, lubricate the O-rings with non-silicon based O-ring lubricant.
Insert the sensor face down into the sensor chamber.
Push the sensor until it bottoms out.
Route the sensor cable the same way the old sensor cable was routed and connect it to the main PCB.
Calibrate the Eagle 2 unit after replacing H2S/CO sensor.

## Slide 44: Replacing the pump
Remove the three Phillips screws that hold the pump to the bottom plate.
Remove the 2-wire connector from main PCB.
Remove pump inlet and outlet tubing, taking care to mark where each tube attaches.
Screws

## Slide 45: Removing the Main PCB
Make sure you are properly grounded.
Remove and note all the connections on the Main PCB.
Carefully lift circuit board out of bottom case.
P/N 57-0102RK-01
www.rkiinstruments.com

## Slide 46: Clock Battery
Replace the clock battery if Time/Date will not set.
If battery is leaking or corroded.
NOTE: Memory battery is soldered into the main board.
P/N 49-1405RK
www.rkiinstruments.com

## Slide 47: Replacing the Display PCB
Remove and replace display PCB if damaged, corroded or if back light is not functioning.
P/N: 57-0103RK
www.rkiinstruments.com

## Slide 48: Photo Ionization Detector
www.rkiinstruments.com
Section 2: Maintenance & PID Cleaning

## Slide 49: PID Maintenance
If the EAGLE 2 is having calibration problems that cleaning the lamp could not solve, the lamp may need to be replaced.
A contaminated or broken electrode stack may also cause calibration problems or PID failure.

## Slide 50: PID Cleaning
Inspect the PID for contamination on the detection window. Contamination will appear as a blue hue
PID Lamp

## Slide 51: Opening the PID
Use tool to compress tabs to remove PID Lamp

## Slide 52: Removing the PID
Grasp lamp with fingers and pull out.
Use finger cots to protect lamp

## Slide 53: PID Cleaning
Use PID lamp cleaning kit 82-0300RK
  Vial of cleaning compound contains a very fine aluminum oxide powder.
  Do not breathe vapor or dust.
  Avoid contact with skin, eyes and clothing.
  Wear suitable PPE when cleaning or handling the lamp
  Keep container closed to prevent water adsorption and contamination.

## Slide 54: PID Cleaning
Open vial of aluminum oxide polishing compound.
Using a clean cotton swab, collect a small amount of compound.
Use cotton swab to polish the PID lamp window.
  Use a circular action applying light pressure
  Never touch the lamp window with fingers!

## Slide 55: PID Cleaning
Polish the PID lamp window until a audible “squeaking” is heard.
  Usually about 15 seconds
Remove residual powder with clean cotton swab

## Slide 56: PID Cleaning
Ensure lamp is completely dry and any visible signs of contamination are removed before refitting.
Reinstall lamp into base.

## Slide 57: Reassemble PID
Reassemble PID and calibrate as needed.

## Slide 58: Common Troubleshooting
www.rkiinstruments.com

## Slide 59: Questions?
www.rkiinstruments.com