Specification for Potable gas monitor

25th June 2014

RIKEN KEIKI CO.,LTD.

Product：Portable Gas Monitor

Model ：GX-6000

Marking

ATEX： II 1G ExiaIICT4 Ga,  Ta = -20 – 50℃

IECEx： ExiaIICT4 Ga,  Ta = -20 – 50℃

【Structure】

<Sensor s and measuring gases>

The GX-6000 can measure maximum six gases with six sensors.

Standard unit measures four gases with four sensors for general combustible gases(LEL), Oxygen(O2), Hydrogen Sulfide(H2S) and Carbon Monoxide(CO).

For other remaining two slots are for Smart Sensors which consist of sensor part and circuit board and are connected with apparatus through digital signal output so various sensors. Three different types of detection principle are applied for Smart Sensors and up to two sensors can be installed into the GX-6000.

Table list of sensors

|  |  | Measuring gas | Detection principle |
| --- | --- | --- | --- |
| Standard four gas |  | Combustible gas (LEL) | Catalytic combustion |
|  | ② | Oxygen (O2) | Galvanic cell |
|  | ③ | Hydrogen Sulfide (H2S) | Electrochemical |
|  | ④ | Carbon Monoxide (CO) | Electrochemical |
| Smart Sensor | ⑤ | VOC | PID |
|  | ⑥ | Toxic gases | Electrochemical |
|  | ⑦ | Carbon Dioxide (CO2)or Combustible gas | Infrared ray (IR) |

<Gas sampling method>

Gas is sampled by a built-in micro pump

<Battery unit>

Either alkaline battery pack “BUD-6000” or lithium-ion battery pack “BUL-6000” can be installed into GX-6000.

Structure of battery unit allows end users to replace it by themselves.

It is supposed to replace the battery unit, alkaline battery, and charge the rechargeable battery at non-hazardous area. Also, Charging BUL-6000 should be done with a specific model, BC-6000 or SDM-6000.

Refer to the figures on the next page.

GX-6000 main unit & Battery unit

Front elevation

GX-6000 main unit

BUD-6000

BUL-6000

BUL-6000

BUD-6000

Rear elevation

GX-6000 main unit

【Schedule for certification】

IECEx and ATEX certification is scheduled to be applied by end of July 2014.

It is expected that IECEx and ATEX certification is approved by end of December 2014.

<Other remarks>

At the end of July 2014, Riken Keiki will supply test samples of plastic case that would not be final version but as cast article. However in September 2014, Riken Keiki will supply the test samples with the final version material.

It is supposed to adjust schedules of the strength tests such as piezo shock test, drop test etc. accordingly.

【Main components and past approved records】

・Sensor

Standard 4 gases

Combustible gas	Sensor only TR	NL/KEM/ExTR07.0057

GX-2009	IECEx KEM 08.0013, KEMA 08ATEX0048

GX-2012GT	IECEx DEK 11.0045, DEKRA 11ATEX0123

Oxygen		GX-2009 	IECEx KEM 08.0013, KEMA 08ATEX0048

GX-2012GT	IECEx DEK 11.0045, DEKRA 11ATEX0123

Hydrogen sulfide	GX-2009 	IECEx KEM 08.0013, KEMA 08ATEX0048

GX-2012GT	IECEx DEK 11.0045, DEKRA 11ATEX0123

Carbon monoxide	GX-2009 	IECEx KEM 08.0013, KEMA 08ATEX0048

GX-2012GT	IECEx DEK 11.0045, DEKRA 11ATEX0123

Smart sensor

PID			Mini PID	IECEx BAS 07.0030U, Baseefa 07ATEX0060U

(manufactured by ION SCIENCE)

Electrochemical	GX-8000	IECEx KEM 10.0038, KEMA 10ATEX0085

Infrared ray (IR)	Lamp only TR	NL/DEK/ExTR12.0033

RX-8000		IECEx DEK 13.0091, DEKRA 13ATEX0228

（Sensor shape differs）

・Micro pump    	GX-2012GT	IECEx DEK 11.0045, DEKRA 11ATEX0123

・Vibration motor		GX-2009	IECEx KEM 08.0013, KEMA 08ATEX0048

・Battery unit single cell	Li-ion: GX-8000	IECEx KEM 10.0038, KEMA 10ATEX0085

DRY: OX-07	IECEx KEM 08.0022, KEMA 08ATEX0126

・Charge circuit/ charger	GX-8000	IECEx KEM 10.0038, KEMA 10ATEX0085

GX-2012GT	IECEx DEK 11.0045, DEKRA 11ATEX0123

・Piezo buzzer		GX-2012GT	IECEx DEK 11.0045, DEKRA 11ATEX0123

・Plastic material  	GX-8000	IECEx KEM 10.0038, KEMA 10ATEX0085

・Panel sheet material	GX-2009 	IECEx KEM 08.0013, KEMA 08ATEX0048

Attached documents

・DIAGRAM FOR I.S. KEEP FOR MODEL GX-6000：E3-6991-5361-10-01K(DRAFT).pdf

・BLOCK DIAGRAM	 : E3-6991-5363-60-01K(DRAFT).pdf

・OUTER STRUCTURE GX-6000(BUL-6000)　　：M3-0009-13-04K.pdf

・OUTER STRUCTURE GX-6000(BUD-6000)　　：M3-0009-13-06K.pdf

・SMART SENSOR　　：M4-0009-08-20K.pdf